## Slide 1
Design of the IOTA Ion Beam Transport System for Nuclear Physics and Medical Applications
A. Zadvornaya | T. Day Goodacre
University of Manchester
Notes: Good morning. I will present how we are using COMSOL Multiphysics to design IOTA, a dedicated offline ion-source optimisation and beam-transport facility at the University of Manchester. The main message is that IOTA is not only a beamline. It is a controlled platform for improving several linked efficiency terms, and COMSOL lets us prioritise the design choices before manufacture and commissioning. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx (uploaded poster)

## Slide 2
Outline
1
Motivation for offline ion source development
2
IOTA objectives and beamline design
3
Integrated COMSOL model
4
Simulations results so far
5
What comes next?
Radioactive isotopes applications, efficiency limitations and online facilities constrains
Purpose, operating principle and layout
Electrostatics, particle tracing, space charge and vacuum
Transmission efficiency, energy spread, and vacuum performance
Further model refinement and validation during commissioning
Notes: CERN-MEDICIS provides a useful example of why a dedicated offline separator matters. It receives irradiated materials, heats them to release the isotopes, ionises them and performs mass separation before collection. Importantly, stable-beam development was used to optimise operation. For samarium, stable tests reached separation efficiencies up to 31 percent, and the optimum temperature was then confirmed during radioactive collections. This is the type of systematic development approach that motivates IOTA, although IOTA will focus strongly on ion-source and transport optimisation in a UK laboratory. [Sources] C. Duchemin et al., CERN-MEDICIS: A Review Since Commissioning in 2017, Frontiers in Medicine 8 (2021) 693682, uploaded PDF, pp. 4 and 8-9

## Slide 3
Radioactive isotopes in cancer treatment and fundamental nuclear physics
CANCER TREATMENT
Targeted alpha therapy (TAT)
Alpha-emitters: Ac-225, At-211, Pb-212
FUNDAMENTAL NUCLEAR PHYSICS
Produce and study
short-lived nuclei far
from stability
Benchmark and
improve nuclear models
A. Gottberg, NIM B, 376 (1) 8-15 2016
TRIUMF | Canada
ONLINE
Target and ion source are connected to the accelerator.
OFFLINE DEVELOPMENT
Supports target and ion-source optimization
Ra-224
L. P. Gaffney et al. 2013 Nature 497 199
Credit: S. Kelley/NIST
Dedicated offline facilities enable systematic ion source development without using valuable online beamtime
F. Maldonado Millan et al, NIM B, 542 (2023) 95 - 98
Notes: CERN-MEDICIS provides a useful example of why a dedicated offline separator matters. It receives irradiated materials, heats them to release the isotopes, ionises them and performs mass separation before collection. Importantly, stable-beam development was used to optimise operation. For samarium, stable tests reached separation efficiencies up to 31 percent, and the optimum temperature was then confirmed during radioactive collections. This is the type of systematic development approach that motivates IOTA, although IOTA will focus strongly on ion-source and transport optimisation in a UK laboratory. [Sources] C. Duchemin et al., CERN-MEDICIS: A Review Since Commissioning in 2017, Frontiers in Medicine 8 (2021) 693682, uploaded PDF, pp. 4 and 8-9

## Slide 4
1
Produce
Irradiate target material
→
→
3
Transfer
Transport from target container to the ion source
→
7
Collect / Measure
Deliver ions to the end station
2
Release
Extraction from heated target
→
5
Transport
Ion-optical transport through the beamline system
→
6
Mass
separate
Select the isotope
of interest
Eff. ~ 10% - 100%
4
Ionize
Create ions in an ion source
Eff. ~ 0.1% - 90%
→
Dedicated R&D is essential to improve isotopes availability, beam intensity, purity, stability and efficiency
Eff.
A. Gottberg, NIM B, 376 (1) 8-15 2016
Y. Martinez Palenzuela et al., Front. Med. 8, 689281 (2021)
CERN-MEDICIS | Switzerland
OFFLINE MASS SEPARATOR
Irradiated material is transported for heating, ionisation, mass separation and collection.
Production and mass separation of radioactive isotopes
Notes: CERN-MEDICIS provides a useful example of why a dedicated offline separator matters. It receives irradiated materials, heats them to release the isotopes, ionises them and performs mass separation before collection. Importantly, stable-beam development was used to optimise operation. For samarium, stable tests reached separation efficiencies up to 31 percent, and the optimum temperature was then confirmed during radioactive collections. This is the type of systematic development approach that motivates IOTA, although IOTA will focus strongly on ion-source and transport optimisation in a UK laboratory. [Sources] C. Duchemin et al., CERN-MEDICIS: A Review Since Commissioning in 2017, Frontiers in Medicine 8 (2021) 693682, uploaded PDF, pp. 4 and 8-9

## Slide 5
Objectives:
Increase the extracted ion current
Improve selectivity: enhance the suppression of long-lived species and develop isotope-specific laser ionisation schemes
Ion source development test stand
Resonance Ionization Laser Ion Source
Resonance Ionisation Mass Spectrometry
Prototype UK-based hot offline mass separator
Infrastructure collaboration with TRIUMF (CA)
Ion source collaboration with The Paul Scherrer Institute (CH)
IOTA: an offline ion source development laboratory
2 kV source bias
Selective laser ionization
Detector station
Einzel lens_1
Einzel lens_2
Deacceleration
Notes: The motivation starts from the radioactive-ion-beam intensity equation. The overall efficiency contains release, transfer, ionisation, transport, separation and any later stages. This means there are many places where targeted development can improve the delivered beam. Large ISOL facilities are oversubscribed and development is constrained by beam time and resources. IOTA provides an offline environment where we can test these steps systematically. In this talk I focus on the transport term, while keeping energy spread and vacuum performance as co-equal design constraints. [Sources] A. Gottberg, Nuclear Instruments and Methods in Physics Research B 376 (2016) 8-15, uploaded PDF, pp. 8-9 C. Duchemin et al., Frontiers in Medicine 8 (2021) 693682, uploaded PDF Slides for Sasha (1).pptx, uploaded background slides

## Slide 6
3
Integrated COMSOL model of the IOTA beamline
1
Electrostatics
Calculate a 3D electric field
›
2
Charged Particle tracing
Track 5,000 Sm⁺ particles
›
Bidirectionally Coupled Particle Tracing
Electrostatics Particle Tracing (Sm⁺: macroparticles, varying release current)
›
4
Free Molecular Flow
Vacuum calculations, including vacuum pumps, gas inflow, and outgassing walls
›
5
Post-processing
in Python
 efficiency, energy and time-of flight (TOF) profiles
Meshing
Goal for the optimization
maximize transport efficiency of the ion beam through the beamline while approaching the 10 eV energy-spread target
Coarser
Extremely fine
Notes: The model is deliberately integrated. We first solve the three-dimensional electrostatic field. Charged Particle Tracing then follows 5,000 samarium ion macroparticles. The particle distribution feeds back into the field through bidirectional space-charge coupling. In parallel, Free Molecular Flow evaluates the proposed pumping arrangement. Python post-processing extracts transport efficiency, the 90 percent energy width and time-of-flight distributions. Iteration is continued until the relative change of a global parameter is below one percent. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx, methodology and Figure 1 (uploaded poster)

## Slide 7
Charged Particle Tracing in Electrostatic Fields (no space charge)
2 kV extraction, 1 uA beam current
94% transport efficiency
13 eV energy spread (centered at 86 eV)
End of beamline:
Source
Maxwell-Boltzman velocity distribution
Uniform angular distribution
2 kV source
0 V: extraction electrode
 high transmission efficiency and small energy spread
Energy distribution
Time-of-flight (TOF)
Notes: The model is deliberately integrated. We first solve the three-dimensional electrostatic field. Charged Particle Tracing then follows 5,000 samarium ion macroparticles. The particle distribution feeds back into the field through bidirectional space-charge coupling. In parallel, Free Molecular Flow evaluates the proposed pumping arrangement. Python post-processing extracts transport efficiency, the 90 percent energy width and time-of-flight distributions. Iteration is continued until the relative change of a global parameter is below one percent. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx, methodology and Figure 1 (uploaded poster)

## Slide 8
Charged Particle Tracing in Electrostatic Fields (space charge included)
2 kV extraction, 1 uA beam current
Space-charge effects
Electrostatics Particle Tracing
(Sm⁺: 5000 macroparticles, varying release current)
Solved iteratively using a bidirectionally coupled particle tracing
Global variable monitored change between successive iterations
Solution was considered as converged when the relative change fell below 1%
30% transport efficiency
27 eV energy spread (centered at 86 eV) at the end of beamline
Smaller transport efficiency and larger energy spread:
Energy distribution
Time-of-flight (TOF)
Notes: The model is deliberately integrated. We first solve the three-dimensional electrostatic field. Charged Particle Tracing then follows 5,000 samarium ion macroparticles. The particle distribution feeds back into the field through bidirectional space-charge coupling. In parallel, Free Molecular Flow evaluates the proposed pumping arrangement. Python post-processing extracts transport efficiency, the 90 percent energy width and time-of-flight distributions. Iteration is continued until the relative change of a global parameter is below one percent. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx, methodology and Figure 1 (uploaded poster)

## Slide 9
Vacuum performance from Free Molecular Flow simulations
Predicted beamline pressure (depending on total gas load):
mbar
Source gas inflow: [(mbar*L)/s]
Walls outgassing rate: [(torr*L)/m^2/s]
Pumping configuration: 2 x 340 [L/s] turbo pumps
Target operating pressure:
Next steps:
Include source temperatures up to  higher outgassing (and source deformations)
Model ion beam interaction with gas
Gas inflow
0 kV: extraction electrode
Paschen curves for helium
Notes: The model is deliberately integrated. We first solve the three-dimensional electrostatic field. Charged Particle Tracing then follows 5,000 samarium ion macroparticles. The particle distribution feeds back into the field through bidirectional space-charge coupling. In parallel, Free Molecular Flow evaluates the proposed pumping arrangement. Python post-processing extracts transport efficiency, the 90 percent energy width and time-of-flight distributions. Iteration is continued until the relative change of a global parameter is below one percent. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx, methodology and Figure 1 (uploaded poster)

## Slide 10
Conclusions
Thank you | Questions
Integrated COMSOL model helps to identify optimal ion optics parameters (geometry and voltages) before commissioning.
The initial IOTA design, developed without space-charge effects, achieved 94% transport efficiency and a 13 eV energy spread at 1 uA current and 2 kV source voltage.
After including space-charge effects, transmission decreased to 30% and the energy spread increased to 27 eV  the next step is to increase the initial acceleration voltage from 2 kV to (5-10) kV
The predicted beamline pressure is mbar, below the maximum allowed of mbar  the next step is to include ion source heating up to and include ion beam interaction with gas.
NEXT: refine model parameters and validate the results during the commissioning
Notes: To conclude, the integrated model predicts high low-current transmission, identifies a strong space-charge penalty at higher current, and quantifies both energy and vacuum performance. The value of COMSOL is not only the final numbers. It is the ability to map coupled design trade-offs before manufacture, and then to update the model using commissioning data. Thank you, and I am happy to take questions. [Sources] AZadornaya-Poster-COMSOL-Conference-2026-Cambridge_2026-09-01.pptx (uploaded poster)